## Slide 1: User Training
Cardiohelp Heart-lung Support System
This document is intended to provide information to an international audience outside of the US. Subject to modifications. All images used are for illustrative purposes only and may not accurately represent the product or its use.

## Slide 2: Table of Contents
Cardiohelp System Overview
3
Connections and Components
5
Disposables: HLS Module Advanced
13
Disposables: Sets with Quadrox-iR
26
Controls
30
Patient Safety Mechanisms
46
Transport
50
Cleaning
54

## Slide 3: Cardiohelp System Overview

## Slide 4: Cardiohelp System Overview
Hardware & Accessories
Cardiohelp Device
Disposable
HLS Set Advanced and other disposable sets
Catheters & Cannulae
HLS Cannulae
Avalon Elite
Refer to the Cardiohelp and HLS Set Advanced IFU. This content is not exhaustive and should not be considered a replacement/substitute.

## Slide 5: Connections and Components

## Slide 6: Cardiohelp Connections and components
Connections and components
Front View
Rear View

## Slide 7: Cardiohelp Connections and components
Front View - Connections
[1] Alarm output (Nurse/Ward call)
[2] Ethernet*
[3] USB port type A – for data export to memory stick
[4] USB port type B – for data recording system
[5] AC power supply
[6] Equipotential Bonding Pin
[7] DC power supply
[8] CAN bus interface*
[9] ECG interface*
[10] Emergency Mode button
1
3
2
4
6
5
7
8
9
10
* Not available
Keep electrical connections covered by protective caps when not in use.
Notes: Discuss Equipotential Bonding Pin and USB Memory Stick.

## Slide 8: Cardiohelp Connections and components
Rear View - Drive and Sensor Connections
[1] External pressure sensors
[2] HLS Internal sensors holder
[3] Clamp*
[4] HLS Internal sensor (pressure and temp)
[5] Drive
[6] Venous probe
[7] Bubble sensor
[8] Fixing pin for HLS Module
[9] External temperature sensors
[10] Unlocking of HLS Module
[11] Level sensor
[12] Flow and bubble sensor
[13] For Battery Calibration
6
5
4
1
10
3
8
2
7
9
11
13
12
* Not available

## Slide 9: Cardiohelp Connections and components
Rear View - Connections
[1] Safety Bar (not a handle!)
[2] Venous probe
[3] Drive

## Slide 10: Cardiohelp Connections and components
  Optical (contactless, spectrometric measurement)
      Hemoglobin Hb
      Hematocrit Hct
      Venous blood saturation SvO2
  Non-invasive, thermopile infrared sensor
      Venous blood temperature TVen
Sensors – Venous Probe
[1] Optical sensors
[2] Connection for venous probe sensor cable
Venous Probes are matched with Cardiohelp devices. They cannot be switched between units.
Methylene blue and other compounds can affect the optical properties of blood and influence SvO2 Measurements.

## Slide 11: Cardiohelp Connections and components
  Arterial Flow-bubble Sensor (FBS)
        3/8” x 3/32”
        Cable length 0.9 m
        Ultrasound detects bubbles ≥ 5mm
  Venous Bubble Sensor (BS) (optional)
        3/8” x 3/32”
        Cable length 1.7 m
        Ultrasound detects bubbles ≥ 5mm
Sensors - Flow and Bubble
Venous Bubble Sensor
Arterial Flow- Bubble Sensor

## Slide 12: Disposables: HLS Module Advanced

## Slide 13: Disposables: HLS Module Advanced
Integrated Centrifugal Pump
    On the blood inlet side of the diffusion membrane oxygenator
Integrated Pressure Sensors
    Venous pressure pVen (pre-pump)
    Internal pressure pInt (pre-membrane)
    Arterial pressure pArt (post-membrane)
Integrated Temperature Sensor
    Optical measurement via infra-red sensor
Integrated measuring cell for venous SvO2, Hb, Hct and temperature.
Overview

## Slide 14: Disposables: HLS Module Advanced
Technical Specifications
 | HLS Set Advanced 5.0 | HLS Set Advanced 7.0
Blood Flow Range (lpm) | 0.5 – 5.0 | 0.5 – 7.0
Membrane surface area | 1.3 m2 | 1.8 m2
Heat Exchanger surface area | 0.3 m2 | 0.4 m2
Volume of Set | 570 ml | 600 ml
Coating |  | 
*Devices with BIOLINE Coating are contraindicated for patients with known hypersensitivity to heparin or patients currently suffering from, or with a history of, type II antibody-related heparin-induced thrombocytopenia (HIT).
*

## Slide 15: Disposables: HLS Module Advanced
Integrated pump on the blood inlet side of the gas-exchanger
Rear View
[1] Gas inlet
[2] De-airing membrane
[3] Centrifugal pump
[4] Water connectors
[5] Venous cell
[6] Connection for internal sensors
[7] Venous Sampling
2
7
6
5
4
3
1

## Slide 16: Disposables: HLS Module Advanced
Blood Flow - Inlet
Blood inlet to the centrifugal pump
Blood in
Blood out

## Slide 17: Disposables: HLS Module Advanced
Venous Measuring Cell
Integrated measuring cell for:
SVO2 - Oxygen Saturation
Hct - Hematocrit
Hb - Hemoglobin
TVen – Venous Temperature

## Slide 18: Disposables: HLS Module Advanced
Pressure Measurements
Pressure sensor post-pump = pressure pre-oxygenator
Venous pressure sensor

## Slide 19: Disposables: HLS Module Advanced
4 unique blood outlets
Blood Flow – Centrifugal pump into oxygenator
1
2
3
4

## Slide 20: Disposables: HLS Module Advanced
Internal Sensors
pven Venous pressure (pre-pump)
pint Internal pressure (pre-membrane)
pArt Arterial pressure (post-pump)
Tart Arterial temperature
Δ p Pressure drop
(calculated from pint - pArt)

## Slide 21: Disposables: HLS Module Advanced
Blood Flow - Oxygenator
De-airing
membrane
Blood inlet
Blood outlet
Luer port

## Slide 22: Disposables: HLS Module Advanced
Oxygenator Function
Polymethylpentene Gas Exchange Hollow Fibers
Plasma-tight—O2 and CO2 gas exchange through diffusion.
  Not compatible with liquid inhalation anesthetics (e.g. isoflurane, ethrane, enflurane).

## Slide 23: Disposables: HLS Module Advanced
Pressure Measurement (cont.)
Arterial pressure sensor and infrared temperature sensor

## Slide 24: Disposables: HLS Module Advanced
HLS Advanced Layout
Venous Cell
Centrifugal Pump
PVen
PInt
PArt
Drainage Cannula
Return Cannula
Δ P
Hb Hct SvO2 TVen
TArt
Gas/Heat Exchange
Gas Exchange
Heater Unit
Water Lines

## Slide 25: Disposables: Sets with Quadrox-iR

## Slide 26: Disposables: HLS Module Advanced
Technical Specifications
 | CIS (Cardiac Intervention Set) | ODP (Organ Donor Perfusion Set)
Catalog Number | CIS 95400 | BE-ODPS 95401 BO-ODPS 95401
Cardiohelp thApp | MECC | 
Typical Place of Use | Cath Lab, Hybrid OR | ICU, ER, OR
Blood Flow Range (lpm) | 0.5 – 7.0 | 
Volume of Set (ml) | ~600 | ~600
Coating | Uncoated | Bioline, Softline or Uncoated
Duration of Use | ≤ 6 hours | 
Quadrox-iR Module
Microporous hollow fibers
Same advantages of HLS Set Advanced: low pressure drop, high O2/CO2 transfer, efficient heat exchange, low prime volume, 1,8 m2 surface area, optional arterial filter.

## Slide 27: Cardiac Intervention Set – CI Set
Application
Display will not be locked after 3 minutes
Place of use: Cath Lab, Hybrid OR
Fields of application:
  Acute cardiopulmonary arrest or severe cardiogenic shock during intervention
  Stand-by or prophylactic support during e.g. high risk PCI
  Bridge to decision

## Slide 28: Cardiac Intervention Set – ODP Set
Application
Level sensor (drib chamber)
Place of use: ICU, ER, OR
Fields of application:
  Cadaver perfusion, donation after cardiac death (DCD)
  Bridge to explant
  Allowed in specific countries

## Slide 29: Controls

## Slide 30: Controls
Control Panel
[1] LED speed indicator
[2] Graphical user interface (GUI)
[3] Battery LED
[4] Power supply LED
[5] Safety button
[6] Zero flow mode LED & button
[7] ON/OFF LED & button
[8] Lock/Unlock LED & button
[9] Rotary knob
1
2
6
5
4
8
7
3
9

## Slide 31: Controls
User Interface – Circulatory and Pulmonary Support
[1] thApp
[2] LPM data
[3] RPM data
[4] Startup screen
[5] Menu screen
[6] Global override
[7] Alarm pause
[8] Power supply status
[9] Tab bar
[10] Parameter display
1
2
6
5
4
8
7
9
10
3

## Slide 32: Controls
Menu Options

## Slide 33: Controls
Flexible Therapy Driven Information

## Slide 34: Controls
Tab Screens – Circulatory and Pulmonary Support

## Slide 35: Controls
  Select any blood parameter (SvO2, Hb or Hct) to enter menu.
  [1] Initialize probe if needed (while attached to safety bar).
Blood Parameters
3
2
a
b
1
  Recalibrate SvO2, Hb and Hct with lab values:
  [2] Save reference values. At the same time, take a blood sample.
  Once the values come back from the laboratory:
    select [3] Recalibrate Values.
    Confirm that the lab value corresponds to the time saved [a].
    Enter lab values in the fields [b]. Confirm.
Notes: Cardiohelp User Manual Section 6.3.1 Recalibrating Blood Parameters

## Slide 36: Controls
Alarm Limit – when exceeded, triggers alarm and Intervention (when active).
Warning Limit – triggers alarm when exceeded.
When Intervention active, exceeding Alarm Limit can increase/reduce pump speed, stop the pump, or go into backflow prevention (depending on the specific parameter).
Warning Limits, Alarm Limits and Intervention Controls
Deactivate Upper Limit: press [+] until [---] is displayed
Deactivate Lower Limit: press [-] until [---] is displayed
Notes: Cardiohelp User Manual Section 6.1 Physiological Alarms and Interventions

## Slide 37: Controls
Calibrate flow sensor with no flow and tubing clamped upstream and downstream.
High-priority alarm when Flow Warning Limit exceeded
Flow Alarm
Notes: Cardiohelp User Manual Section 6.1 Physiological Alarms and Interventions

## Slide 38: Controls
Backflow Prevention
  Cardiohelp monitors blood flow and detects and responds to backflow of blood.
  Backflow of -0.1 l/min is the fixed Alarm Limit:
      1 second of backflow  medium-priority alarm
      6 seconds of backflow  high-priority alarm and zero flow mode activates
  Backflow Prevention can also be triggered by a Venous Bubble Intervention.
  Zero Flow Mode
  CARDIOHELP targets a flow of 0 l/min by controlling the pump accordingly (both forward and reverse)
  The Zero Flow Mode can be activated manually or automatically by Backflow Prevention.
  Activate Zero Flow Mode: hold down the “Safety” button and press the “Zero Flow Mode” button .
  Deactivate Zero Flow Mode: increase RPM high enough to ensure forward flow. press the “Zero Flow Mode” button .
Flow Monitoring
Notes: Backflow Prevention: Cardiohelp User Manual Section 6.2.1 Flow Monitoring
Zero Flow Mode: Cardiohelp User Manual Section 3.3.2 Keypad
Can be demonstrated by applying pressure from the Emergency Prime Line

## Slide 39: Controls
Pump Speed Monitoring (only in LPM Mode)
Setting an appropriate RPM alarm always recommended.
RPM Alarms/Intervention only available in LPM Mode. RPM Intervention not enabled when switching to LPM mode.
When Alarm Limit exceeded, intervention will reduce/increase pump speed and cause med-priority alarm.
Low-priority alarm when Alarm Limit exceeded (Intervention off)
Notes: Cardiohelp User Manual Section 6.1 Physiological Alarms and Interventions

## Slide 40: Controls
Interventions - Pressure
Zero sensor when HLS Set is empty.
Low-priority alarm when Warning Limit exceeded.
When exceeding the Alarm Limit by <10 mmHg with intervention on, Cardiohelp automatically controls the pump speed and generates a medium-priority alarm. Exceeding the Alarm Limit by >10 mmHg will stop the pump and cause high-priority alarm.
Low-priority alarm when intervention off and Alarm Limit exceeded.
Only lower limits available for pVen. Only upper limits available for pArt und pInt.
Notes: Cardiohelp User Manual Section 6.1 Physiological Alarms and Interventions

## Slide 41: Controls
Microbubbles ≥ 5 mm triggers alarm. Smaller bubbles may be detected.
Bubble Monitoring
High-priority alarm when bubble detected. Arterial FBS intervention stops the pump and Venous BS intervention activates backflow prevention.
Low-priority alarm when bubble detected (Intervention off).
To reset the bubble alarm.
Notes: Cardiohelp User Manual Section 6.1 Physiological Alarms and Interventions

## Slide 42: Controls
Other Warning Limits generate low-priority alarms when exceeded:
Pressure Drop (Δp)
Temperature (TArt,TVen)
Blood Parameters (SvO2, Hb and Hct)
Other Alarms

## Slide 43: Controls
Battery Power
[1] charging and power supply status
[2], [3], [5] Battery Status
[4] Remaining battery time
[6] Date of last battery change
Battery life: at least 90 minutes (with fully charged batteries)
Unplugging external power: audible alarm and pop-up message.
Battery <20%: medium priority alarm
Battery <10%: high-priority alarm
Notes: Cardiohelp User Manual Section 4.9.2 Displaying Battery Status

## Slide 44: Cardiohelp Set-up & Priming
See Cardiohelp System Set-up Guide:
Preparation & Installation
Priming the System
Connecting to the Patient
Starting Perfusion
Ending Perfusion
Notes: https://getinge.training/all-courses/cardiohelp-outside-us/

## Slide 45: Patient Safety Mechanisms

## Slide 46: Patient Safety Mechanisms
Clamp red and blue lines.
If applicable, mount the Emergency Drive on the Cardiohelp or on the Sprinter Cart Emergency Drive Holder.
Open the Safety Bar
Disconnect: a) Integrated Sensor Cable, b) Venous Probe and c) Transport Guard (if applicable).
If needed, temporarily stop the flow of the Heater Unit and disconnect Water Lines while the HLS Set is being moved.
Disconnect the HLS Set from the Cardiohelp.
[1] Align the top of the HLS Set into the emergency drive puncturing the protective film. [2] Press the locking switch while swinging the bottom part of the HLS Set into place.
Release clamps and start turning the hand crank clockwise at the desired speed.
Transfer to new Cardiohelp asap.
Emergency Drive
Provides manual drive to the HLS Set if the Cardiohelp fails.
  Always keep Emergency Drive accessible at all times.
Refer to the Cardiohelp IFU Section 6.4.4 Using the Emergency Drive with the Disposable HLS Retainer and HLS Set IFU Section 8.3 Replacing the Set
Notes: Cardiohelp User Manual Section 6.4.4 Using the Emergency Drive with the Disposable HLS Retainer
Difference in Drive attachment and E-drive attachment. This is a redundancy in case the three plastic fixation tabs on the module get damaged.

## Slide 47: Patient Safety Mechanisms
Emergency Mode
      Activate Emergency Mode: press Emergency Mode button waiting for 3 beeps and release.
      Note: alarms are also deactivated (not mentioned in the pop-up message)
      Exit Emergency Mode: switch off Cardiohelp by pressing the Emergency Mode button releasing after 10 beeps and a long tone.
Safe perfusion even with a damaged display or controls
Notes: Cardiohelp User Manual Section 6.4.3 Using the Emergency Mode

## Slide 48: Patient Safety Mechanisms
To remove air from the blue line before it has reached the oxygenator.
Clamp Red Line and reduce pump speed to 0 rpm (display will read 0, but continue to turn knob counter-clockwise).
Clamp Blue Line and the Venous Cannula.
Prime the Emergency Priming Line.
Attach Emergency Priming Line to dialysis lock ensuring air-free connection.
Disconnect the connection between the venous cannula and the Blue Line.
Remove air by releasing clamp from the Blue Line and push priming solution retrograde into the Blue Line.
Clamp the Emergency Prime Line.
Form an air-free connection between the de-aired Blue Line and the venous cannula.
Remove clamp from the Red Line.
Reset Venous Bubble Sensor (if applicable).
Restart perfusion.
Remove Emergency Priming Line.
Emergency Priming Line
See HLS Set IFU Section 8.1 Using the Emergency Priming Line for full instructions.
Notes: HLS Set IFU Section 7.1 Using the Emergency Priming Line.

## Slide 49: Transport

## Slide 50: Transport
Transport Accessories
Mobile Holder (HKH 8860) includes Protection Cover
Base Plate (HKH 8850)
Wall Holder (HKH 8820)
Trolley Transport System
Sprinter Cart XL
Notes: Cardiohelp User Manual Section 4.1.3 Attaching the CARDIOHELP to the Mobile Holder HKH 8860
All info contained in IFU “Patient Transport Accessories for Patient Transport” is also repeated in the Cardiohelp IFU

## Slide 51: Transport
Transport Accessories
Transport Guard
Provides extra securement for disposables during transport. For inter-hospital transport and intra-hospital transport without a Sprinter Cart.
Transport Case
Stores Emergency Drive and Transport Guard when not in use. Emergency drive must be stored in case during inter-hospital transport.
Notes: Cardiohelp User Manual Section 4.1 Installing Accessories for Patient Transportation

## Slide 52: Transport
  Disconnect non-essential connections (if applicable: nurse call, USB cable, heater water lines, etc.)
  Check function of transport equipment
    Mount Transport Guard onto Cardiohelp
    Transport Emergency Drive with patient in Transport Case
    Bring protective (rain) cover
  Transfer gas supply (gas tank)
  Unplug from electrical outlet and check batteries. Recommend leaving power cable plugged into Cardiohelp, but use yellow plug cover if it is removed.
  Mount Cardiohelp on Transport Accessories including Mobile Holder.
Preparing for Inter-Hospital Transport
Notes: Cardiohelp User Manual Section 5.6.2 Check Before Every Use in Transportation

## Slide 53: Cleaning

## Slide 54: Cleaning
Clean and disinfect after each use.
Only use approved cleaning agents and disinfectants.
Cleaning the Venous Probe
    Detach push-button for better cleaning.
    Only use lint-free cloth and water to clean the Venous Probe’s sensors
Refer to the Cardiohelp IFU Section 10 Cleaning and Disinfection
Notes: Cardiohelp User Manual Section 10 Cleaning and Disinfection

## Slide 55
Questions?
This document is intended to provide information to an international audience outside of the US.
The products in this brochure may be pending regulatory approvals to be marketed in your country. Contact your Getinge representative for more information.
Cardiohelp, Avalon Elite, Bioline Coating are a registered Trademark by Maquet Cardiopulmonary GmbH
Manufacturer · Maquet Cardiopulmonary GmbH · Kehler Str. 31 · 76437 Rastatt · Germany · +49 7222 932-0